## Slide 1: DEVELOPMENT OF A RP-HPLC METHOD FOR THE DETERMINATION OF METFORMIN IN HUMAN PLASMA
KISCOMS 5th International Congress of Medical Sciences 8-10 May 2015 Prishtina, Republic of Kosovo
Eva troja

## Slide 2
Determination of metformin in biological fluids
(human plasma, urine, etc)
Challenge to conventional chromatographic methods

## Slide 3
Metformin
Oral antidiabetic drug
Small molecule with high polarity
(impossible to maintain by RP columns)
pKa1=2.8 and pKa2=11.5

## Slide 4
Determination of Metformin in plasma
There are many methods to determine Metformin in plasma :
LC-MS/MS
GC
Capillary electrophoresis
Cation-exchange HPLC
Normal phase chromatography
RP-HPLC (specific column or common columns using ion-pairing agent)

## Slide 5
The purpose of the study
To developed a method using a conventional column

## Slide 6: Ion-pairing chromatography

## Slide 7
Sample preparation techniques – Advantages and Limitations
Protein precipitation or dilution
Easy
Not specific
Cost efficient
Extraction is necessary because:
Provides a purity of biological material to be injected
Enables analyte’s concentration
Solid phase extraction
Requires skilled operator
Specific
Less cost efficient when compared to protein precipitation or dilution methods
Liquid-liquid extraction
Requires skilled user
Specific
Less cost efficient when compared to protein precipitation or dilution methods

## Slide 8
What do we want from our method ?
To be simple
To use conventional column
To be selective
To be accurate
To be rapid
To comply the purpose (study of BE from plasmatic data)

## Slide 9
Materials and reagents
Metformin HCl
ACN (HPLC grade)
SDS, H3PO4 (analytical grade)
Millipore water (Direct Q-UV-Ultrapure® water system)
NYLON membrane filters type HNWP (47mm, 0.45µm)

## Slide 10
Instruments/Chromatographic system
Chromatographic system in reverse phase Agilent 1200, having in-built binary pump, UV/Vis DAD (Agilent 1200) using CWS software (ChemStation) install in PC.
LiChrocart ® 100 RP 18 (125mm x 4.0mm i.d., 5µm particle size) C18.
Centrifuge
Analytical balance
Vortex
pH-meter
Ultrasonic cleaner

## Slide 11
Chromatographic conditions
Mobile phase ACN:sodium phosphate buffer 10 mmol pH=6.0, 0.3%SDS (32.5:67.5)
Mobile phase was filtered and degased for 30 minutes before use.
The flow rate 1.25 ml/min, dedection was carry out in 236 nm
Column’s temperature 50°C
Manual injector 20µL
The retention time of metformin was 4 minutes
Method was linear in the range 50-1600 ng/ml.

## Slide 12
  Calibration standards and quality control plasma samples
  Stock solution Metformin in buffer (500 µg/ml)
  Working standards (1-50 µg/ml) obtained by diluting the stock solution in milliwater
  The solutions were freshly prepared for each analytical run
  Blood samples were centrifugated at 3500 rpm/min for 10 min
  Protein precipitation and wash with dichloromethane

## Slide 13
Development and validation of bioanalytical method
The main analytical parameters that are used for validation of bioanalytical method are :
Linearity
Precision
Accuracy
Limit of dedection
Limit of quantification
Selectivity
Specificity
Robustness

## Slide 14
Linearity
Concentration of analyte– absorbance
Calibration curve of metformin in plasma
For research use only. Not for use in diagnostic procedures.
Fig 1. HPLC-UV chromatograms obtained for blank plasma spiked with metformin 50 ng/ml (a), 100 ng/ml (b), 300 ng/ml (c), 800 ng/ml (c), 1500 ng/ml (d), 1600 ng/ml (f).

## Slide 15
Metformin Calibration Curves
a) b)
Fig 2. Calibration curves of metformin in plasma within-day (a) and between days (b).
Chart: ng/mL
| 50 | 200 | 400 | 1600
 | 0.068 | 0.33 | 0.69 | 2.8
Chart: ng/ml
| 50 | 200 | 400 | 1600
 | 0.0705 | 0.305 | 0.67 | 2.8

## Slide 16
Metformin in Plasma – Method Performance
Nominal concentration (ng/ml) | Calculated concentration (ng/ml) | Within-day Mean ± SD (n=6) | Precision % | Accuracy %
Low QC (300) | 312.2 | 312.2±13.09 | 4.2 | 4.1
Medium QC (800) | 819.7 | 819.7±23.06 | 2.8 | 2.5
High QC (1500) | 1591.8 | 1591.8±23.06 | 1.4 | 6.1
Nominal concentration (ng/ml) | Calculated concentration (ng/ml) | Between days Mean ± SD (n=21) | Precision % | Accuracy %
Low QC (300) | 304.0831 | 304±13.78 | 4.5 | 1.36
Medium QC (800) | 791.8740 | 791.8±38.7 | 4.9 | 1.02
High QC (1500) | 1547.92 | 1547.92±51.97 | 3.4 | 3.19

## Slide 17
RECOVERY
Nominal concentration (ng/ml) (n=6) | Recovery % | Precision %
QC – L (300) | 95.3 | 4.68
QC-M (800) | 98 | 5.7
QC-H (1500) | 100.7 | 4.38

## Slide 18
Limit of Dedection and Limit of Quantification (LOD and LOQ)
LOD = 3.3 SD/S and LOQ = 10 S.D/S, where S.D is LOD = 3.3 SD/S and LOQ = 10 S.D/S, where S.D is the residual standard deviation of regression line and S is the slope of the line.
LOD = 18 ng/ml
LOQ = 50 ng/ml
In the range 50-1600 ng/ml

## Slide 19
Selectivity and specificity
a)
b)
Fig 3. Spectra of Metformin in 236 nm(a) and HPLC-UV chromatogram of blank human plasma (b)

## Slide 20
Robustness
Small changes in mobile phase organic strength
and in pH by ± 2% have no significant effect on chromatographic resolution.

## Slide 21
Metformin in Plasma- Data Examples
The method was used to determined metformin concentration in plasma samples from 20 healthy volunteers, after oral administration of a tablet containing 850 mg of metformin.
5 samples analysed in 20 minutes
For research use only. Not for use in diagnostic procedure.
Fig. 4. HPLC-UV chromatograms of plasma sample from a healthy volunteer a- 1 h (c=1322.4 ng/ml) b- 2 h (c=1708.1ng/ml) c- 3 h (c=1873.4ng/ml) after oral administration 850 mg of metformin.

## Slide 22
Application in Pharmacokinetics and Bioequivalence Studies
Fig 5. Mean plasma concentrations of metformin (ng/ml) versus time (hr) profile for 20 volunteers after administration of Metforminë Profarma Sh.a 850 mg Tablets and Glucophage® 850 mg Tablets.
Chart: Hours
| 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 8 | 10 | 12 | 14
Metformin | 0 | 894.5496 | 1316.2192 | 1488.464 | 1575.5398 | 1635.1661 | 1574.4138 | 1346.1983 | 1110.5893 | 873.2067 | 574.5896 | 391.0634 | 239.4535 | 154.051
Glucophage | 0 | 746.3094 | 1319.1693 | 1426.2615 | 1513.2048 | 1543.0959 | 1516.7757 | 1376.0215 | 1142.8096 | 909.5273 | 608.5397 | 393.5752 | 251.31 | 169.298

## Slide 23
Conclusions
Method was simple,
rapid
selective
accurate
use conventional column
comply the purpose (study of BE from plasmatic data )

## Slide 24
Acknowledgements

## Slide 25
THANK YOU FOR YOUR ATTENTION !